AutoISF Fork —Brief HCL Operations Manual

This is the plain-language summary to the full AutoISF Operations Manual -- same section numbering, same headings, same settings/reference tables, but each subsection here is one short paragraph explaining what it does and why it matters, instead of the full technical detail. Use this to get oriented quickly; use the full manual [or Plain Language manual] when you need the actual mechanism, thresholds, or code references.

This manual concentrates on the operational changes that most distinguish this AutoISF fork from stock AAPS and that are most likely to affect daily use:

# AIM  of the program. [a personal use example of ideas]

The aim is to make the requirement for intervention in Hybrid Closed Looping very minimal, given still the expectation of at least intermittent user interference for bolus calculation.

This represents a lesser goal than “Full Closed Loop (FCL)”, for which more focussed instructions are available relation to AutoISF, or likewise the “Boost” Variant AAPS programme that can be aimed at no- bolus as well.

Although the internal code automations do switch profiles and set temp targets these can be set individually as well if necessary and particularly if there is a concern about overall excess  strength causing hypoglycemia or various conditions [such as steroids, hormones and so on] interfering in the last few days or weeks,

It is simple enough to switch the “high and low” profiles used by the program.  Also, setting “MJ” types from List1  [or States] sets either 3, 2 , or 1 day of more gentle criteria for changing profiles, setting Steroid profiles [first set Kotlin MJ and Steroid buttons] from List1 sets more aggressive changes during the day. [note: personal use emphasis, only used by one person, with FSL2+ and pods]

# Introduction: Most significant new features  [↑ TOC]

# Screenshots

Real device screenshots (Virtual, 24 Aug 2026). Small thumbnails floated beside the relevant Introduction paragraphs link here -- click one to jump to its full-size copy, or just scroll through them below in one place.

GUI display changes.img_overview_stacked

This graph is always the primary time-aligned display Also reachable via a direct TT-code Secondary graphs (graphs 1-4). Graph 5 is a dedicated main-sized comparison panel Gesture and remote triggers. A basal-area long-press cycles the main graph's display preset and controls the quick visibility of empirical carb line 1. An IOB-area long-press toggles SMB labels, resets the basal display preset to normal   [↑ TOC]

Bolus wizard: split-bolus + Walking soon   [↑ TOC]

- Full description of the bolus calculator changes; split and delayed bolus, safer calculation;

Changes to bolus calculator with addition,s modificationeg

Initial screen and live calculation., carbs, protein, fat . Protein and fat are not included in the immediate total

Manual split-bolus if  bolus over maximum bolus, remainder split. Delayed bolus for low bgl criteria.

Import settings, with Keep-current checkboxes   [↑ TOC]

- Settings Import Features — Especially With States.

Four “Keep current…” checkboxes not present in stock AAPS

Automation States are force-disabled after every import, requiring confirm

Import restores states wholesale — where “states I never made myself” can come from

List 1 -- Direct AutoISF settings   [↑ TOC]

- Lists 1 and 2: full descriptions; rapid access to settings

Double-tap the IOB area to open List 1. Double-tap the basal-rate area to open List 2

The outer list shows setting names ; for setting preferences and some functions, with confirmation.

On AAPSClient preferences are not changed locally, so confirmation creates the matching five-minute custom temporary-target code for the pump to receive, apply and cancel.

Steroid Kotlin button on/off confirmation   [↑ TOC]

MJ Kotlin buttons on/off confirmation   [↑ TOC]

List 2 -- Direct actions   [↑ TOC]

- AIV table popup, automatic backup with logs and settings immediately and 6 hourly
Long-pressing the ISF/history control opens the AIV history popupEach row presents grouped glucose/raw-delta, ISF-factor, insulin/carb/HP, request/delivery and activity/status fields

Each row joins the nearest available context to an AIV timestamp and presents grouped glucose/raw-delta, ISF-factor, insulin/carb/HP, request/delivery and activity/status fields Manual button press userentries or export logs and  Automatic /export ; all export AIV csv, text, setting txt, userentries.csv, logarithmic zipTOC_main

17. 	New carer/ parent helpers

# In Virtual pump sync settings, option to import insulin dosing as well , but not smbs, onlt manual bolus with the calculator entry

# Still may be slow to pick up the NS entry, but better if Virtual pump is being used to check usefulness of changed settings on its own loop..

# Directs to the secondary settings [+/- separate website if NS site for virtual is different {eg Juggluco Raw data upload}] to accept the bolus entries/ carbs/ sensor and pump change inputs [therapy events]End   [↑ TOC]

# Introduction: Most significant new features  [↑ TOC]

This is the plain-language companion to the full AutoISF Operations Manual -- same section numbering, same headings, same settings/reference tables, but each subsection here is one short paragraph explaining what it does and why it matters, instead of the full technical detail. Use this to get oriented quickly; use the full manual [or Plain Language manual] when you need the actual mechanism, thresholds, or code references.

This manual concentrates on the operational changes that most distinguish this AutoISF fork from stock AAPS and that are most likely to affect daily use:

- Full description of the bolus calculator changes; split and delayed bolus, safer calculation;

Changes to bolus calculator with addition,s modificationeg

Initial screen and live calculation., carbs, protein, fat . Protein and fat are not included in the immediate total

Manual split-bolus if  bolus over maximum bolus, remainder split. Delayed bolus for low bgl criteria.

Automatic delayed bolus at 50% or State recent-50. This is distinct from the manual split.

- Lists 1 and 2: full descriptions; rapid access to settings

Double-tap the IOB area to open List 1. Double-tap the basal-rate area to open List 2

The outer list shows setting names ; for setting preferences and some functions, with confirmation.

On AAPSClient preferences are not changed locally, so confirmation creates the matching five-minute custom temporary-target code for the pump to receive, apply and cancel.

- Settings Import Features — Especially With States.

Four “Keep current…” checkboxes not present in stock AAPS

Automation States are force-disabled after every import, requiring confirm

Import restores states wholesale — where “states I never made myself” can come from

- AIV table popup, automatic backup with logs and settings immediately and 6 hourly
Long-pressing the ISF/history control opens the AIV history popupEach row presents grouped glucose/raw-delta, ISF-factor, insulin/carb/HP, request/delivery and activity/status fields

Each row joins the nearest available context to an AIV timestamp and presents grouped glucose/raw-delta, ISF-factor, insulin/carb/HP, request/delivery and activity/status fields

Manual button press userentries or export logs and  Automatic /export ; all export AIV csv, text, setting txt, userentries.csv, logarithmic zip

- Mechanism For Ignoring Or Invoking Native Automations; new ones now self-healing for the coded automations' own states.

his fork re-implements roughly 100+ of what used to be native AAPS Automation-tab triggers/actions directly as hand-coded Kotlin blocks

each native Automation-tab event's title against that registry every time it decides whether to actually run a native event

cf explanation of automation toggles gor MJ, Steroid buttons, installation check of automations code turned on for aISF running

- Tier 3 UAM Boost criteria, output limits and duration.

Its Virtual-Pump-only restriction was removed 2026-08-27; it can now activate on a real pump when the master switch is on.

Tier3 from Boost variant program with different protection; testing only

- BolusGiven strong-boost effects.

BolusGiven strong boost. This is separate from Tier 3. sets SMB delivery ratio to Mild Boost base + 0.03 It uses delivery change

- BolusGivenMild and Mild Failsafe [tduration and stack limits] behavior.

Mild Failsafe applies similar effects BolusGiven/Mild automations run first and can change smb_delivery_ratio, profile percentage, target and weights. AutoISF weighs BG acceleration, post-meal rise, and how long BG has been stable differently depending on the situation -- each has a normal (“resting”) weight and a stronger (“boosted”) one that kicks in when conditions call for extra correction.

- Steroid-management initiation and escalation.

Steroids (e.g. prednisone) sharply and progressively raise insulin resistance for as long as they're active. • GUI: “Steroid Kotlin button” setting Also reachable via a direct TT-code Each step requires the previous step's Automation State + profile match to still hold [or set directly from List1

- The Mounjaro three-day reduced-insulin cycle.

MJ means Mounjaro in this fork  The MJ controls implement a reduced-insulin state cycle for the period after a Mounjaro dose, when insulin requirements may be lower and hypoglycaemia risk may be greater Also reachable via a direct TT-code The start button is offered when Automation State MJ is NOMJremains

- Hypoglycaemia warnings and protective overrides.	
Hypoglycaemia warning causes and outcomes.

Hypoglycaemia warnings do not automatically mean “cancel MJ and restore 100%”. GentleHypoRisk, AlarmHypo1, AlarmHypo2, the 50% profile machinery, and related HP/HP2 checks can lower acceleration weighting. SMS may also be sent

- GUI display changes.

This graph is always the primary time-aligned display Also reachable via a direct TT-code Secondary graphs (graphs 1-4). Graph 5 is a dedicated main-sized comparison panel Gesture and remote triggers. A basal-area long-press cycles the main graph's display preset and controls the quick visibility of empirical carb line 1. An IOB-area long-press toggles SMB labels, resets the basal display preset to normal

- VirtualPseudoWizard: background real-bolus correction (Virtual Pump, testing only).
It places an actual BS.Type.NORMAL bolus, computed by running the same constrained BolusWizard used by the manual calculator, but entirely in the background — no calculator screen is ever shown This graph is always the primary time-aligned display

- Automation States in use: which states this fork’s coded automations read/write and why.

Appendix F — Automation States In Use  ; A reference for every named state this fork’s coded automations

- Fast-rise SMB tapering: the tiered cascade, the separate late taper, and what the AIV export actually shows.
a set of checks for excess insulin SMB size, scaling them down

- Fast-rise SMB tapering: the tiered cascade, the separate late taper, and what the AIV export actually shows.
a set of checks for excess insulin SMB size, scaling them down

- 17. New carer/ parent helpers

In Virtual pump sync settings, option to import insulin dosing as well , but not smbs, only manual bolus with the calculator entry ; better if Virtual pump is being used to check usefulness of changed settings on its own loop..

Contents   [↑ TOC]

Introduction: Most significant new features — p.1

1. Requirements for “Automation States” — p.13

a. It has a master on/off switch, and things silently do nothing when it's off — p.13

b. Fixed 2026-08-19: this is now self-healing for the coded automations' own states — p.13

c. A second, related gap: declaring the value list still isn't the same as the state having a value — p.13

2. Settings Import Features — Especially With States — p.13

a. Automation States are force-disabled after every import — by design, not a bug — p.14

b. Import restores states wholesale — where “states I never made myself” can come from — p.14

c. Four “Keep current…” checkboxes not present in stock AAPS — p.14

3. Settings That Materially Affect Day-to-Day Function — p.14

a. The SMB delivery brake — p.14

b. Tier 3 UAM Boost: criteria, output, and duration — p.14

c. The weight system's Resting/Boosted pairs — p.14

d. Duration (dura) modifications to the AutoISF algorithm — p.14

e. TDD factor / sensitivity toggles — p.15

f. The overnight safety-guard architecture — p.15

4. Setting: If Steroids Are Needed — p.15

a. Turning it on — p.15

b. The escalation ladder — p.15

5. Mounjaro (MJ): three-day reduced-insulin cycle — p.15

a. Starting the cycle — p.15

b. Day-by-day state progression — p.16

c. What changes while MJ remains active — p.16

d. Hypoglycaemia warnings and protective overrides — p.16

e. Automatic exit and manual restoration — p.16

6. Mechanism For Ignoring Or Invoking Native Automations — p.16

a. The reconciliation mechanism — p.16

b. The suppression only applies while AutoISF is actually the active algorithm — p.16

7. Other Idiosyncrasies Worth Knowing — p.16

a. CarePortal note codes as a compact diagnostic language — p.17

b. MJ as the fork's other major state machine — p.17

c. The AIV/UKFcheck export pipeline has more moving parts than it looks — p.17

d. A concrete regression pattern worth remembering — p.17

8. Manual controls and AIV history — p.17

a. Bolus calculator: screen, calculation and follow-up dosing — p.17

b. Lists 1 and 2: mechanism, current status and confirmation — p.17

c. AIV table: popup population and backup/export behavior — p.17

d. Coded profile roles (Standard/Low): assignment, re-picking, and why it matters — p.19

e. Steroid escalation as a third coded role (added 2026-08-23) — p.19

9. GUI display changes — p.19

a. Graph types and the triggers for each — p.19

b. Colour changes driven by AutoISF input — p.19

c. Information shown on each graph — p.19

d. What the different carbohydrate graphs mean — p.19

10. Factors affecting progress and hypoglycaemia outcomes — p.19

a. Pod age and pump-type qualification — p.19

b. Projected HP at bolus-calculator initiation — p.20

c. Effects of 50recent, a 50% profile and a hypo temporary target on progress — p.20

d. Hypoglycaemia warning causes and outcomes — p.20

11. UKF: formulas, roles and initial comparison assessment — p.20

a. The UKF variants and where each is used — p.20

b. Formula and full explanation of each — p.20

c. Progress of the initial UKF-check assessment and earliest-change question — p.20

d. UKF1 for live AutoISF dosing (toggle, added 2026-08-23) — p.20

12. Battery 1% automation and profile recovery — p.20

a. Battery1pc: action at 1% or below — p.20

b. BatteryOver1pc: restoration after recovery — p.20

13. SensorAge: aging/new-sensor Libre calibration compensation — p.20

a. What it compensates for and how the tiers are built — p.21

b. The six age tiers and their symmetry — p.21

c. Two independent enable switches — p.21

d. Entry/exit gating and interaction with LowRaw24Cal — p.21

e. Related sensor/pod-age reminders (notification-only) — p.21

14. VirtualPseudoWizard: background real-bolus correction (Virtual Pump, testing only) — p.21

a. Why it is not a real carb treatment — p.21

b. Entry gating — p.21

c. Calculation and delivery mechanics — p.22

d. Outcome reporting — p.22

15. Fast-Rise SMB Tapering: the Tiered Cascade and the Separate Late Taper — p.22

a. Where this sits, and the entry gate — p.22

b. The tier ladder: Delta/SDelta bands, BG sub-branches, and multipliers — p.22

c. Other branches inside FAST RISE HANDLING: Higher BG, early-morning guard, twilight hours — p.22

d. The preceding sibling branches, including Sensor Glitch3 — p.22

e. The separate late taper and cumulative SMB caps — p.22

f. What actually reaches the AIV "FastRise" column — p.22

g. Libre-calibration slope compensation (added 2026-08-22) — p.23

Appendix A — Installation and explanatory information — p.25

Appendix B — Changelog — p.25

Appendix D — Alphabetical CarePortal Note Code Reference — p.25

Appendix E — List 1 and List 2 Temporary-Target Command Catalogue — p.31

Appendix F — Automation States In Use — p.32

Appendix G — Current Live Settings Snapshot — p.37

16. Screenshots — p.1

17. New carer/parent helpers — p.9

# 1. Requirements for “Automation States”    [↑ TOC]

“Automation States” is a generic named state-machine store (its own plugin, its own “States” tab in the UI) shared between native Automation-tab triggers/actions and this fork's own coded automations in Kotlin. Two things make it easy to trip over:

## a. It has a master on/off switch, and things silently do nothing when it's off  [↑ TOC]

Automation States are a shared memory the coded automations use to track things like whether steroids are active or a low just happened. There's one master switch for the whole system -- if it's off, every automation that depends on a state just quietly does nothing, rather than erroring.

## b. Fixed 2026-08-19: this is now self-healing for the coded automations' own states  [↑ TOC]

If the app needs a state value that was never registered (say, after an update added a new one), it used to be able to crash. Now it automatically registers anything missing the first time it starts, so this fixes itself instead of needing a manual settings edit.

## c. A second, related gap: declaring the value list still isn't the same as the state having a value  [↑ TOC]

Declaring which values a state is allowed to hold isn't the same as that state actually holding one yet -- a freshly registered state can still be genuinely empty until something writes to it for the first time.

Also fixed 2026-08-19, same day: the bootstrap now also assigns a safe default CURRENT value, but only for the 5 states where there's real evidence of what “neutral/off” means (never guessed): MJ → NOMJremains, MJstate → MJoff, Steroids → Steroids Off, LowBG → NO50rec, Profile → AllOK. It deliberately leaves Steps, Sleeping, BGLstate, and query_got_up declare-only — no coded automation in this file was ever seen writing the first two (something else owns them), BGLstate has only one enumerated value with no observed “off” counterpart (defaulting to it would wrongly assert a low just happened on every fresh device), and query_got_up is a one-shot signal, not a toggle.

The default is only applied when the state currently has no value at all — it can never overwrite something you (or the import) already set.

# 2. Settings Import Features — Especially With States  [↑ TOC]

Settings import in this fork is substantially more elaborate than stock AndroidAPS. The underlying mechanism is still “clear every SharedPreferences key, then write the file's contents back in” (ImportExportPrefsImpl.doImportSharedPreferences()) — but several fork-specific safeguards sit around that blunt operation.

## a. Automation States are force-disabled after every import — by design, not a bug  [↑ TOC]

Restoring a settings backup automatically switches Automation States off -- on purpose, not a bug -- since an old backup's states might not match what's really happening on the phone right now (e.g. it might think steroids are still active when they aren't). You confirm and turn it back on manually after checking the states still make sense.

## b. Import restores states wholesale — where “states I never made myself” can come from  [↑ TOC]

Importing a settings file overwrites every state value at once with whatever that file contains -- not just numeric preferences. So a state you never touched yourself can still change on import, if it happened to be recorded in that particular backup.

## c. Four “Keep current…” checkboxes not present in stock AAPS  [↑ TOC]

Four checkboxes not found in stock AAPS let you protect specific state values from being overwritten during an import -- useful when you want most of an old backup restored but need today's real current state kept for a few particular things -- including one specifically to keep the current pump driver, its settings, and any active pod/session state, so importing an older backup can't overwrite a running pod.

# 3. Settings That Materially Affect Day-to-Day Function  [↑ TOC]

The full settings tree (~65 entries) is documented setting-by-setting in the companion “AutoISF Settings Reference”. This section instead highlights the handful whose day-to-day real-world effect is outsized relative to how easy they are to change by accident, or how non-obvious their interaction with something else is.

## a. The SMB delivery brake  [↑ TOC]

A safety ceiling on how much insulin SMB can deliver in a short window, regardless of what the dosing math would otherwise suggest -- the backstop that limits how aggressive one burst of automated dosing is allowed to get.

## b. Tier 3 UAM Boost: criteria, output, and duration  [↑ TOC]

An extra dosing boost that raises how much insulin the loop is allowed to give during a genuinely fast, unannounced rise. Its Virtual-Pump-only restriction was removed 2026-08-27, after a full day of real Virtual-Pump validation -- it can now activate on a real pump too when the master switch is on. Trigger (changed 2026-08-24): no independent gate of its own any more — Tier 3's sole trigger is now BolusGivenMild's own entry gate (bmildBasicCriteriaMet), the same rise-confirmation BolusGivenMild's real dosing already relies on.

Sizing checks that remain: IOB below its own ceiling, boost scale under 3, BG above 4.4mmol, and remaining IOB allowance above zero. Removed 2026-08-26: two forward-prediction checks (eventualBG above target, insulinReq above zero) that real data showed made the combined gate almost unreachable — an existing IOB on board routinely makes that forward prediction disagree with BolusGivenMild's own immediate rise-detection at exactly the moments that matter.

Diagnostic logging (added 2026-08-25) now reports Tier 3's calculated candidate against BolusGivenMild's ordinary amount every cycle the gate is entered, win or lose — not just on a win. Two further observation-only additions (2026-08-24) never touch dosing: a reference-formula comparison (CarePortal note T3AcceISF) and a fast-carb-rebound check (T3FastCarb), both using a different, looser gate than the real trigger, so they can fire independently of it.

## c. The weight system's Resting/Boosted pairs  [↑ TOC]

AutoISF weighs BG acceleration, post-meal rise, and how long BG has been stable differently depending on the situation -- each has a normal (“resting”) weight and a stronger (“boosted”) one that kicks in when conditions call for extra correction.

## d. Duration (dura) modifications to the AutoISF algorithm  [↑ TOC]

The “duration” part of AutoISF's sensitivity calculation -- how long BG has been stable -- has fork-specific tuning on top of the stock algorithm, changing how much a long stable stretch is allowed to loosen sensitivity.

## e. TDD factor / sensitivity toggles  [↑ TOC]

Settings that let your recent total daily insulin dose influence sensitivity calculations, plus toggles to turn that influence on or off.

## f. The overnight safety-guard architecture  [↑ TOC]

Several independent guards work together specifically overnight: hard ceilings on how far a value like IOB threshold can shift, escalating anti-stacking limits on SMB delivered close together, and forced switches to a gentler profile late at night. All added after two real overnight low-BG incidents, each one catching a different failure mode the others missed.

# 4. Setting: If Steroids Are Needed  [↑ TOC]

Steroids (e.g. prednisone) sharply and progressively raise insulin resistance for as long as they're active. This fork has a dedicated, fully worked-out mechanism for it — the most complete example of a “medical-event ladder” anywhere in the codebase, and a reasonable template if a similar mechanism is ever built for another condition (see §5).

## a. Turning it on  [↑ TOC]

A dedicated on/off switch and profile-escalation path for when you're taking steroids, which sharply raise insulin resistance -- turning it on shifts AutoISF onto a stronger baseline immediately.

## b. The escalation ladder  [↑ TOC]

Steroid insulin needs typically build up over days, so there's a five-step ladder (110%→130%→150%→190%→250%) you move through manually as resistance increases, each step switching to a correspondingly stronger profile and raising the IOB ceiling.

| Step | Profile switched to | iobTH % | bgAccel weight |
| --- | --- | --- | --- |
| START (110%) | Current Profile110Real | 75 | 0.70 |
| INCREASE (130%) | Steroid Profile130 | 78 | 0.70 |
| INCREASE (150%) | Steroid Profile150 | 81 | 0.70 |
| INCREASE (190%) | Current Profile190Real | 84 | 0.70 |
| INCREASE (250%) | Current ProfileReal250 | 88 | 0.70 |
| TURN OFF | Current ProfileReal | 71 | 0.71 |

# 5. Mounjaro (MJ): three-day reduced-insulin cycle  [↑ TOC]

MJ means Mounjaro in this fork. The MJ controls implement a reduced-insulin state cycle for the period after a Mounjaro dose, when insulin requirements may be lower and hypoglycaemia risk may be greater. The Overview button describes the intended operating period as three days of the configured low-profile strategy; the implementation uses named states and clock-time transitions rather than an exact 72-hour countdown.

## a. Starting the cycle  [↑ TOC]

A three-day reduced-insulin automation for the days around a Mounjaro (tirzepatide) injection, since insulin needs drop sharply for a short window afterward -- you trigger it manually when you inject.

## b. Day-by-day state progression  [↑ TOC]

The three days aren't treated identically -- the automation tracks which day of the cycle you're on and adjusts its own thresholds and duration, and profile choices as the cycle progresses back toward normal.

## c. What changes while MJ remains active  [↑ TOC]

While an MJ cycle is running, several other automations (profile switches, overnight guards) deliberately defer to MJ's own reduced-insulin logic instead of fighting it.

## d. Hypoglycaemia warnings and protective overrides  [↑ TOC]

MJ's reduced-insulin window genuinely raises hypo risk, so extra warning and override paths exist specifically for BG behaviour seen during an active MJ cycle.

## e. Automatic exit and manual restoration  [↑ TOC]

The cycle ends on its own after the third day, reverting profiles back to normal -- but can also be cancelled manually early if needed, or reset to any day, and hypo warnings may reinstate day3 setting MJ3.

# 6. Mechanism For Ignoring Or Invoking Native Automations  [↑ TOC]

This fork re-implements roughly 100+ of what used to be native AAPS Automation-tab triggers/actions directly as hand-coded Kotlin blocks inside OpenAPSAutoISFPlugin.kt (the “coded automations” layer described in CLAUDE.md), so they can react every 1-minute loop cycle instead of on the native automation engine's slower tick. Both systems — native Automation-tab events and coded automations — remain fully present in the app at the same time, which creates a real risk of the same logical action firing twice (once from each system) unless something reconciles them.

## a. The reconciliation mechanism  [↑ TOC]

Since this fork reimplements 100+ native AAPS Automation-tab triggers as hand-coded Kotlin logic, it needs to stop the ORIGINAL native automation from also firing and double-acting. A registry compares every native automation's title against the coded equivalents and suppresses the native one whenever a coded version already covers it.

| Match | Meaning | Effect on the native event |
| --- | --- | --- |
| EXACT | normalized title equals a coded key exactly | Always suppressed — the coded equivalent is assumed to be running every cycle instead, no user choice involved |
| CLOSE | one normalized string contains the other, without being exactly equal | Suppressed only if the user hasn't explicitly accepted it via the review popup (below); a prior “accept” decision (persisted per title) lets it run |
| NONE | no relationship to any coded key | Never touched — runs exactly as stock AAPS would |

## b. The suppression only applies while AutoISF is actually the active algorithm  [↑ TOC]

This suppression only happens while this fork's AutoISF algorithm is the one actually running the loop -- switch to a different algorithm and native automations are free to run normally again.

# 7. Other Idiosyncrasies Worth Knowing  [↑ TOC]

## a. CarePortal note codes as a compact diagnostic language  [↑ TOC]

Every coded automation leaves a short, consistent CarePortal note (like “BsUp” or “DelOff”) each time it fires, forming a compact log you can read back afterward to see exactly what happened and when. See Appendix D for the full glossary of what each code means [for example].

| Code family | Meaning |
| --- | --- |
| AVLs / AVLf | AIV local export succeeded / failed |
| ACEs / ACEf | combined-export (combined<Name>.txt) rebuild succeeded / failed |
| UKCs / UKCn / UKCf | UKFcheck diagnostic export found matches / ran fine but found nothing / genuinely failed (UKCf can ONLY come from an exception — never means “no findings”) |
| StB On / StB Off | Steroid Kotlin button feature toggled on/off via its TT-code |
| SteroidsON / Steroids130…250 / SteroidsOff | each step of the steroid ladder (§5) |

## b. MJ as the fork's other major state machine  [↑ TOC]

Alongside the general Automation States system, MJ has its own separate day-tracking state machine dedicated to its three-day cycle.

## c. The AIV/UKFcheck export pipeline has more moving parts than it looks  [↑ TOC]

The AIV history export (and the UKF comparison tool) pull from several different sources and reconstruction paths rather than one simple table -- worth knowing before assuming every exported column is a raw, directly-stored value.

## d. A concrete regression pattern worth remembering  [↑ TOC]

A real bug already happened once from changing an automation's ACTION value without re-checking its own condition for clearing/re-arming against that new value. Worth deliberately checking both sides together whenever retuning a threshold.

# 8. Manual controls and AIV history  [↑ TOC]

## a. Bolus calculator: screen, calculation and follow-up dosing  [↑ TOC]

The Wizard/QuickWizard bolus calculator handles carbs, protein and fat, offers a manual split-bolus option when a dose exceeds your configured maximum, and can automatically deliver a delayed top-up dose at 50% profile or after a recent low -- plus a “Walking soon” checkbox that halves the immediate dose for anticipated activity.

## b. Lists 1 and 2: mechanism, current status and confirmation  [↑ TOC]

Two quick-access settings menus -- double-tap the IOB icon for List 1, the basal-rate icon for List 2 -- give direct toggles, nudges and actions without navigating the full Settings screen. On a real pump or Virtual these apply immediately; on AAPSClient they relay via a short temporary-target code instead.

## c. AIV table: popup population and backup/export behavior  [↑ TOC]

Long-pressing the ISF/history area opens a scrollable history table built from AIV records, and separately kicks off a full export (CSV, text, settings, UserEntries, logs) -- both on that open and automatically every six hours regardless, and on Manual button press userentries or export logs.

| Command [can set directly on Virtual and Live] | Relay TT (mmol/L) | Current effect |
| --- | --- | --- |
| SMBdel base + mild-Bst | 5.002 / 5.004 | Decrease/increase both SMB-delivery baseline and Mild Boost ratio by 0.01. |
| Toggle Libre sensor adjustment | 5.006 | Toggle aging-sensor/Libre adjustment on or off. |
| Toggle boost automations | 5.008 | Toggle the whole boost-automation group on or off. |
| pp ISF weight - normal | 5.012 / 5.014 | Decrease/increase normal postprandial ISF weight by 0.01. |
| Acceleration ISF weight - normal | 5.016 / 5.018 | Decrease/increase normal acceleration ISF weight by 0.05. |
| Duration ISF weight - normal | 5.022 / 5.024 | Decrease/increase normal duration ISF weight by 0.1. |
| Libre slope - original | 5.026 / 5.028 | Decrease/increase the baseline Libre calibration slope by 0.01. |
| Libre offset - original | 5.032 / 5.034 | Decrease/increase the baseline Libre calibration offset by 0.05. |
| SMB offset | 5.036 / 5.038 | Decrease/increase the hard SMB-offset override by 0.1. |
| Clean graph view | 5.042 | Hide SMB dose labels/arrows and show the plain solid-green view. |
| Wizard bolus percentage | 5.046 / 5.048 | Decrease/increase the Wizard percentage by five percentage points. |
| Mild Boost ratio | 5.052 / 5.054 | Decrease/increase Mild Boost ratio alone by 0.01. |
| pp ISF weight - high | 5.056 / 5.058 | Decrease/increase boosted postprandial ISF weight by 0.01. |
| Acceleration ISF weight - high | 5.062 / 5.064 | Decrease/increase boosted acceleration ISF weight by 0.01. |
| Higher-ISF-range weight | 5.068 / 5.070 | Decrease/increase high-BG ISF weight by 0.1. |
| Peak insulin time | 5.074 / 5.076 | Decrease/increase insulin peak by five minutes. |
| AutoISF maximum - low BG | 5.080 / 5.082 | Decrease/increase the low-BG AutoISF maximum by 0.1. |
| AutoISF maximum - normal | 5.086 / 5.088 | Decrease/increase the normal AutoISF maximum by 0.1. |
| TOD offset 00:00-02:00 | 5.092 / 5.094 | Decrease/increase this time-of-day offset by 0.1. |
| TOD offset 02:00-04:00 | 5.098 / 5.100 | Decrease/increase this time-of-day offset by 0.1. |
| TOD offset 04:00-06:00 | 5.104 / 5.106 | Decrease/increase this time-of-day offset by 0.1. |
| TOD offset 06:00-09:00 | 5.110 / 5.112 | Decrease/increase this time-of-day offset by 0.1. |
| TOD offset 09:00-12:00 | 5.116 / 5.118 | Decrease/increase this time-of-day offset by 0.1. |
| TOD offset 12:00-18:00 | 5.122 / 5.124 | Decrease/increase this time-of-day offset by 0.1. |
| TOD offset 18:00-22:00 | 5.128 / 5.130 | Decrease/increase this time-of-day offset by 0.1. |
| TOD offset 22:00-00:00 | 5.134 / 5.136 | Decrease/increase this time-of-day offset by 0.1. |
| Toggle Graph2 carb-model curve | 5.138 | Toggle the empirical/theoretical carb-model curve display. |
| Cloud logs upload | 5.140 | Zip and send logs to configured cloud storage, otherwise use email. |
| Toggle Graph5 | 5.142 | Toggle the Graph5 comparison panel. |
| MJ state manual override | 5.144 / 5.146 | Set MJ to NOMJremains or MJ3. |
| Profile manual override | 5.148 / 5.150 | Select the configured Standard or Low profile. |
| Toggle UKFset1 live comparison | 5.152 | Toggle the Virtual-Pump, non-Client live UKFset1 comparison; hidden elsewhere. |
| Run SensorAge code | 5.156 | Toggle the sensor-age code on or off. |
| Command [can set directly on Virtual and Live] | Relay TT (mmol/L) | Current effect |
| MJ start | 5.158 | Start the confirmed Mounjaro cycle. |
| MJ restore | 5.160 | Run the confirmed Mounjaro restoration action. |
| Steroid start | 5.162 | Start the steroid-management sequence. |
| Toggle MJ Kotlin buttons | 5.164 | Show/hide the MJ buttons. |
| Toggle Steroid Kotlin button | 5.166 | Show/hide the steroid button. |
| Steroid 110 to 130 | 5.168 | Run the confirmed steroid escalation to 130. |
| Steroid 130 to 150 | 5.170 | Run the confirmed steroid escalation to 150. |
| Steroid 150 to 190 | 5.172 | Run the confirmed steroid escalation to 190. |
| Steroid 190 to 250 | 5.174 | Run the confirmed steroid escalation to 250. |
| Steroids OFF | 5.176 | Run the confirmed steroid turn-off action. |
| AnyDesk restart | 5.178 | Client writes an ADesk Note and uses a five-minute relay TT only if no real TT is active; otherwise Note-only transport preserves the real TT. |
| AnyDesk restart - local test | 5.180 | Local command test through the real receiver; no NS round-trip or relay TT. |
| Graph: UKF1 | No relay TT | Local display toggle plus optional Libre slope/offset calibration. |
| Graph: UKF2 | No relay TT | Local display toggle; calibration is already upstream and remains shown as on. |
| Graph: UKF3 | No relay TT | Local display toggle plus optional Libre slope/offset calibration. |
| Graph: Graph5 panel | No relay TT | Local panel toggle; optional BGL-only mode hides insulin activity and the three carb series. |

## d. Coded profile roles (Standard/Low): assignment, re-picking, and why it matters  [↑ TOC]

AutoISF's automations don't hardcode which profile counts as “Standard” or “Low” -- you assign your actual profiles to those two roles once, and every automation reads that assignment live. Re-picking later updates every single automation at once, with no code change needed.

Two real incidents from live AIV/UserEntries data, both traced the same day, that clarify what this system does and doesn't cover. First: a real BasalUp ("BsUp") firing on 23 Aug correctly switched to "Current ProfileReal" rather than the user's newly-created "Current Profile100" -- not a bug, just timing: the Standard role hadn't been re-picked to point at the new profile yet at that exact cycle, confirmed by tracking autoisf_standard_profile_name across every settings snapshot that day, and the very next coded-automation profile switch (the reinstated "Switch to Standard/Low profile now" List 1 row, "PrSt" note) correctly used the new name once the re-pick had landed.

Second, more important: a LATER switch back to "Current ProfileReal" traced to a source literally named "button1" in the UserEntries log -- not "AutoISF code-based profile switch" or "AutoISF: reset to 100%", which is what every one of THIS file's ported automations always logs. "button1" is a native AAPS Automation-tab trigger, configured entirely outside this codebase (Settings -> Automation), with the old profile name typed directly into its own Profile Switch action. No amount of re-picking the coded roles touches it -- it has to be edited by hand in the Automation tab itself.

## e. Steroid escalation as a third coded role (added 2026-08-23)  [↑ TOC]

The six steroid strength levels are now assignable roles too, the same way Standard/Low already were -- renaming or replacing a steroid profile no longer needs touching any code.

# 9. GUI display changes  [↑ TOC]

## a. Graph types and the triggers for each  [↑ TOC]

The main graph and up to four secondary graphs can each show a different combination of series (IOB, COB, treatments, and more); a fifth dedicated comparison panel (Graph 5) is also available.

## b. Colour changes driven by AutoISF input  [↑ TOC]

Graph points are tinted by whichever AutoISF factor (acceleration, BG, postprandial, or duration) is furthest from neutral at that moment, so you can see at a glance what's driving the current sensitivity adjustment -- low-glucose points always show the low-BG colour regardless.

## c. Information shown on each graph  [↑ TOC]

Each graph panel carries its own fixed extra rows beyond whatever series you've selected -- delta figures, hypo-prediction values, step counts, and similar.

## d. What the different carbohydrate graphs mean  [↑ TOC]

Several carb-related lines exist and mean genuinely different things -- entered carbs vs. COB vs. a theoretical absorption curve vs. the deviation-inferred “unannounced carbs” estimate -- worth not conflating them.

# 10. Factors affecting progress and hypoglycaemia outcomes  [↑ TOC]

## a. Pod age and pump-type qualification  [↑ TOC]

Some automations only apply once enough time has passed since the last pod/cannula change. The patch-pump-only restriction on this was removed 2026-08-27 -- infusion-site age is now treated the same way regardless of pump type, though the specific hour thresholds involved were originally tuned against Omnipod's own wear cycle.

## b. Projected HP at bolus-calculator initiation  [↑ TOC]

The bolus calculator runs its own hypo-prediction safety check every time it recalculates, trimming a proposed dose if the projected low afterward looks too close.

## c. Effects of 50recent, a 50% profile and a hypo temporary target on progress  [↑ TOC]

A recent-low flag, an active 50% profile, and a hypo temporary target are related but distinct things -- each can independently reduce dosing, and they don't all mean the same thing.

## d. Hypoglycaemia warning causes and outcomes  [↑ TOC]

Hypo warnings can be triggered by several different, unrelated code paths -- not one single BG threshold.

# 11. UKF: formulas, roles and initial comparison assessment  [↑ TOC]

## a. The UKF variants and where each is used  [↑ TOC]

“UKF1/UKF2/UKF3” aren't three unrelated things -- they're different combinations of Kalman filtering and Libre-specific smoothing, built for different comparison purposes.

## b. Formula and full explanation of each  [↑ TOC]

The exact mathematical formula behind each variant is in the full manual -- skip this unless you need the underlying math.

## c. Progress of the initial UKF-check assessment and earliest-change question  [↑ TOC]

Real-data testing suggests UKF1 reacts to a genuine BG change faster than the standard LibreSpecial smoothing -- but this is still being validated against more real episodes before being trusted for anything beyond comparison. See the next section for the live-dosing toggle this finding led to.

## d. UKF1 for live AutoISF dosing (toggle, added 2026-08-23)  [↑ TOC]

An opt-in, Virtual-pump-only toggle that swaps AutoISF's own dosing-relevant delta/acceleration numbers for the faster-reacting UKF1 signal instead of the usual LibreSpecial-smoothed one. Because it reacts faster, today's gate thresholds may trip a little earlier than they were tuned for -- a compensation setting exists to correct for that once enough real data justifies a value.

# 12. Battery 1% automation and profile recovery  [↑ TOC]

## a. Battery1pc: action at 1% or below  [↑ TOC]

If battery drops to 1% or below on a real pump, the app automatically switches to the configured safety profile (a dedicated setting, no longer a hardcoded profile name) and sends its alerts regardless of whether that switch actually succeeds. A separate, non-blocking check (ProfileRoleSanityCheck) periodically warns if Standard, Low, or the safety profile don't match a real profile.

## b. BatteryOver1pc: restoration after recovery  [↑ TOC]

Once battery recovers above 1% and you're confirmed still on that same safety profile, your normal profile is restored automatically.

# 13. SensorAge: aging/new-sensor Libre calibration compensation  [↑ TOC]

SensorAge (OldSensorAdj in code) overrides the Libre calibration slope/offset (FslCalSlope/FslCalOffset) during known-unreliable stretches of a sensor’s life, tapering back to the user’s configured baseline everywhere else. Unlike most of this fork’s automations it has no dosing effect of its own — it only reshapes the calibrated glucose value that everything else, including dosing, is computed from. Unlike the reminder cluster in §13e below, the core tier-compensation block carries no Virtual-Pump gate at all: it applies identically regardless of pump type.

## a. What it compensates for and how the tiers are built  [↑ TOC]

Libre sensors read less reliably right at the very start of their life, and again toward the very end -- this automation temporarily adjusts the calibration slope/offset during those windows, tapering back to your normal baseline outside them.

## b. The six age tiers and their symmetry  [↑ TOC]

Three tiers cover a brand-new sensor's first three days; three more mirror them for the sensor's last three days before expiry -- each pulling the calibration progressively further from baseline the more extreme the window.

## c. Two independent enable switches  [↑ TOC]

A master switch turns the whole mechanism on or off; a second, narrower switch inside it specifically gates the aging (11-15 day) tiers -- both need to be on for any tier to actually apply. The master switch now has a Settings-screen home too (Settings -> AutoISF settings), not just List 1.

## d. Entry/exit gating and interaction with LowRaw24Cal  [↑ TOC]

A separate automation (LowRaw24Cal) takes priority when it's active, temporarily overriding SensorAge entirely. SensorAge itself only ever applies when a genuinely unreliable-BGL pattern has actually been seen recently (or at least a recent spike) -- not just because the calendar says the sensor is old.

Real cross-device discrepancy found the same day, still not fully explained. Comparing settings snapshots at the same moment: Client and the real pump-driving device ("Live") both showed old_sensor_adj_active=true and a correctly tiered live FslCalSlope (0.70, stepped down from a 0.72 baseline), while Virtual showed old_sensor_adj_active=false and FslCalSlope stuck at the raw, never-adjusted 0.72 baseline -- despite identical toggle states (both enable switches on) on all three.

Live's own data later showed this ISN'T permanently broken -- a later snapshot from Live showed the SAME stuck-at-baseline pattern Virtual had shown earlier -- suggesting the underlying tier-entry condition (recentHighBglSeen/sensorAgeDays/the tier-boundary match) is genuinely flickering cycle-to-cycle on whichever device, not cleanly broken on one device only. The new diagnostic logging (this paragraph, above) is what's needed to catch a live failure with actual field values -- not yet captured, since the logging hadn't been built onto any device as of this writing.

## e. Related sensor/pod-age reminders (notification-only)  [↑ TOC]

Separate, notification-only reminders exist for pod/sensor age that don't touch dosing at all -- they just nudge you.

# 14. VirtualPseudoWizard: background real-bolus correction (Virtual Pump, testing only)  [↑ TOC]

VirtualPseudoWizard is different in kind from Tier 3 UAM Boost, BolusGiven, and BolusGivenMild (§3b), which only ever strengthen the SMB delivery ratio. It places an actual BS.Type.NORMAL bolus, computed by running the same constrained BolusWizard used by the manual calculator, but entirely in the background — no calculator screen is ever shown, and no user confirmation occurs. It genuinely delivers dosing -- a real bolus via the same command queue as any other bolus -- but only ever on Virtual Pump, never a real one; it is hard-gated twice specifically so it cannot fire on a real pump: once in its eligibility check, and again immediately before the bolus is queued.

## a. Why it is not a real carb treatment  [↑ TOC]

This automation places a real bolus using a made-up “carbs” number (10g/15g/20g, tiered by BG), purely to size a correction dose sensibly through the normal Wizard math -- it isn't recording that carbs were actually eaten.

## b. Entry gating  [↑ TOC]

Only fires on Virtual Pump, and only under a fairly narrow, specific set of conditions (very high BG, a fast confirmed rise, a sustained pattern) meant to catch a real runaway high that other automations missed.

## c. Calculation and delivery mechanics  [↑ TOC]

Runs the normal bolus-calculation math in the background and delivers the result automatically, with no user confirmation dialog.

## d. Outcome reporting  [↑ TOC]

Logs a clear success/failure CarePortal note either way, so it's visible afterward whether it actually delivered.

# 15. Fast-Rise SMB Tapering: the Tiered Cascade and the Separate Late Taper  [↑ TOC]

## a. Where this sits, and the entry gate  [↑ TOC]

A tiered ladder that decides how much to trim (or not trim) an SMB dose during a fast rise, based on how fast BG is climbing and by how much.

As of 2026-08-22, rawDelta1Mgdl — the single-minute *raw, unsmoothed* delta — was removed from every gate in this cascade, and the code now carries an explanatory comment at the point where the old check used to sit. The removal was driven by real AIV data from two same-day sustained-rise events (01:50–01:58 and 02:50–02:55): at every single sampled minute of both events, rawDelta1Mgdl read negative or flat (as low as -0.56 mmol) even while Delta, SDelta and rawDelta5Mgdl all clearly confirmed a genuine, sustained rise the whole time.

A single unsmoothed one-minute raw reading is simply too noise-prone to be a reliable confirmation on top of signals that already agree — requiring it was silently blocking the entire cascade during exactly the kind of gradual sustained rise it exists to catch. rawDelta5Mgdl (the 5-minute raw delta) was kept, since it passed cleanly at both real events and still gives a raw-signal check independent of the smoothed AAPS delta.

Checking the current source confirms the removal is complete: rawDelta1Mgdl now appears only as a function-parameter declaration (default sentinel 9999.0, meaning "no data supplied") and in this explanatory comment — it is not referenced in any condition anywhere in the file any more. aapsDelta1Mgdl, by contrast, remains part of the entry gate (and several sibling branches' gates) throughout.

## b. The tier ladder: Delta/SDelta bands, BG sub-branches, and multipliers  [↑ TOC]

Different combinations of how fast BG is rising and what BG level it's rising from select a different multiplier applied to the normal SMB size -- steeper or higher rises get less trimming.

## c. Other branches inside FAST RISE HANDLING: Higher BG, early-morning guard, twilight hours  [↑ TOC]

A few additional special-case branches exist for particularly high BG, early-morning hours, and twilight-hour patterns.

## d. The preceding sibling branches, including Sensor Glitch3  [↑ TOC]

Earlier-checked branches, including one for a specific sensor-glitch pattern, take priority over the main tier ladder whenever they match.

## e. The separate late taper and cumulative SMB caps  [↑ TOC]

Beyond the tier ladder itself, a separate later-stage taper and two cumulative SMB-amount caps (10-minute and 30-minute) provide additional protection against stacking too much insulin in a short window, regardless of which tier fired.

## f. What actually reaches the AIV "FastRise" column  [↑ TOC]

The exported “FastRise” column reflects exactly which tier/branch fired and its multiplier that cycle, reconstructed by pattern-matching the logged reason text -- it isn't a raw stored value.

One further wrinkle worth flagging for anyone reading exported FastRise values from before 2026-08-22: for four of the tiered-ladder branches — mild fast rise’s three BG sub-branches (0.9/0.85/0.75) and the early fast-rise tier (0.7) — the actual multiplier had been loosened from an earlier value via an inline "was X" comment, and the *first* reason-text line for each was updated to match the new multiplier, but the "CHANGED SIZE X.XXX for ... fast rise X.XXX" label’s own numeral was not — so any AIV export taken before 2026-08-22 shows the *pre-loosening* value (0.855, 0.806, 0.707, 0.608) rather than the multiplier genuinely applied that cycle (0.9, 0.85, 0.75, 0.7 respectively) for those four branches.

The moderate tier’s two upper-BG branches had the opposite mismatch — their labels (0.602, 0.653) already tracked their current 0.6/0.65 multipliers correctly, but the *other*, non-label reason-text line in those two branches was stale (still reading "* 0.8" and "* 0.7"). All six were corrected 2026-08-22 so every logged numeral now matches what’s actually applied — but if you’re reconstructing historical FastRise values from an export taken before that date, the pre-fix numbers above are what you’ll actually see in the file, not the current code’s behavior.

## g. Libre-calibration slope compensation (added 2026-08-22)  [↑ TOC]

The whole cascade's mmol thresholds are compensated for how far your live Libre calibration slope has drifted from the 0.75 baseline they were originally tuned against -- so the same real-world rate of rise triggers the same tier regardless of calibration drift. This doesn't extend to basal rate or SMB delivery ratio, since neither is a glucose-kinetics measurement.

`

# Appendix A — Installation and explanatory information  [↑ TOC]

This fork builds as four separate Android apps (“product flavors”), each with its own application ID, its own icon, and its own independent SharedPreferences store. They install side by side on the same device without conflicting — a real device commonly runs two or three of these at once (e.g. the main pump-driving app plus one or two Client mirrors on other phones/tablets).

Because each flavor is a genuinely separate app, a settings export/import done inside one flavor never touches another flavor's preferences, even on the same physical device — “import into AAPS” and “import into AAPSClient” are two unrelated operations.

## a. Version tag & build identification  [↑ TOC]

Every console/history line and the on-device log stamp the running build as the last 13 characters of Config.VERSION_NAME (e.g. “aisf321UK_654”), derived automatically from the actual APK version rather than a hand-maintained literal. Known caveat: on the aapsclient/aapsclient2 flavors, VERSION_NAME gets an extra “-aapsclient”/“-aapsclient2” suffix appended, which shifts the fixed 13-character window — the printed tag can look slightly off on Client builds specifically. Not yet fixed as of this writing.

## b. Per-device backup/export folder naming  [↑ TOC]

The optional “GeneralPatientName” setting scopes where this device's AIV/log exports land — both on-device folder structure (Documents/AAPS/aapsLogs/<Name>/) and the PC-side backup tree pulled off the phone (aiv_<Name>, logs_<Name>, aapsLogs\<Name>\output\). Leaving it blank falls back to a single shared, unscoped folder — fine for one device, ambiguous the moment a second device is added. Set a distinct name per physical/virtual device before its exports start accumulating, since retroactively renaming only affects future exports, not the ones already written under the old name.

# Appendix B — Changelog  [↑ TOC]

Concise dated index of significant changes. Detailed explanations have been integrated into the relevant topical sections so each feature is described where it is used.

### 2026-08-19  [↑ TOC]

Automation-state bootstrapping was made self-healing and given evidence-based neutral defaults. Full operational detail is now incorporated into Section 1.

### 2026-08-22  [↑ TOC]

Bolus and Quick Wizard controls, Fast-Rise signal gating, and FastRise reporting were updated. The current behaviour and historical-export cautions are incorporated into Sections 8 and 15.

### 2026-08-23  [↑ TOC]

Tier 3 safeguards, List controls, coded profile roles, UKF1 options, SensorAge diagnostics, graph labels, battery recovery, and LowBG recovery were updated. Details are incorporated into their corresponding topical sections and Appendix F.

# Appendix D — Alphabetical CarePortal Note Code Reference  [↑ TOC]

Compiled directly from every addCarePortalNote()/addGraphAnnouncement() call site in OpenAPSAutoISFPlugin.kt (current-code audit dated 2026-08-21; cross-checked, not superseded, by this manual’s own prose above). On/Off toggle pairs are merged into one entry; codes built with a numeric or block-identifying suffix (e.g. "MJoff-N", "compactSettingNote" prefixes like "D<val>") are collapsed to their base code, with the suffix’s meaning explained in the entry itself. The free-text "VirtualPseudoWizard SUCCESS/FAILED" notes (§7b’s sibling addition) are intentionally excluded — they are not part of this compact-code language.

• 130Off — HighPP130Off: exits a 130%/110% profile boost (BGL settled with no TT, falling with a TT active, or simply no TT while still boosted) and restores the standard profile.

• 50ff-N (N=1–6) — PP50.Off: exits the 50%-profile hypo-protection state once BGL has recovered; suffix identifies which of 6 recovery blocks matched (daytime stabilisation, fast rise with verified cannula, treated-hypo carbs+bolus, IOB resolved, overnight recovery, or the acce-weight-gate catch-all).

• 50pcTT — "50pc makes5.7": while on the 50% profile with BGL dropping under 5.0mmol and still falling, sets a 150-minute 5.7mmol safety temp target.

• 50Rec — 50SetRecent: flags LowBG=50recent whenever the profile drops back to 50% while previously marked recovered, so downstream low-BG logic knows a 50%-profile episode is (again) active.

• A1 — Test-automation ports "MJ4"/"MJ5" (direct code-test triggers, fired via a manually-driven MJ automation state): advances the MJ state to NOMJremains (MJ5 additionally forces the standard profile for 30 min).

• __________A4 — AlarmHypo2's first graph-announcement marker (TE.asAnnouncement mechanism, renders on the main graph): logs a hypo-alarm event — BGL ≤4.3mmol, or ≤5.5mmol with high recent activity, or a hypo-prediction match.

• AC<val> — AcceWeightDownTT/UpTT: remote ±0.05 nudge (via a manually-set 5.016/5.018mmol temp target) on the acce-ISF-weight baseline (ApsAutoIsfBgAccelWeightNormal); suffix is the new value.

• Ac50 — ActivityProf50%: drops to a 50% profile for 180 minutes while the manually-set 6.8mmol "Activity" temp target is active and BGL is stable/falling.

• Acce — AcceUp0.5: raises the acce-ISF weight to its boosted level (ApsAutoIsfBgAccelWeightHigh) when it's currently low, BGL is stable/rising, no TT is active, and MJ is clear or steroids are on.

• AcLT<role> / AcNS<role> — AnyDesk remote-restart receipt notes (role = C/V/R for Client/Virtual Pump/Real pump): logs that this device accepted an AnyDesk-restart command, either a same-device local test (AcLT) or one relayed via the secondary-NS "ADesk" Note channel (AcNS); does not itself confirm Tasker/AnyDesk actually completed.

• ACrv / PArv / PPrv — PpWeightRevertUnder8_5: automatically reverts a boosted ppISF weight and/or acce weight back to baseline once BG < 8.5mmol, activity is high, or no genuine recent high (raw Libre >12.0mmol/48h) was seen; the code picks which of the two settings (or both, PArv) was reverted.

• ActOff — "ActivityOff": exit criteria for ActivityProf50%'s 50% profile drop — restores the standard profile, acce 0.35, iobTH 70% once the Activity TT ends/BGL stabilises or a bolus was just given.

• ActTToff1 — Activity TT reversal branch "TT6.0New1" (6.5–7.6mmol Activity TT band): cancels the TT on an active-meal recovery (high/rising BG with COB and IOB headroom).

• ActTToff2 — Activity TT reversal branch "TT6.0New2": cancels the same TT on a quiet/settled recovery (moderate BG, very low IOB, flat-or-rising).

• AcTT<role> — AnyDeskRestartActionTT: second, redundant AnyDesk-restart trigger channel via a dedicated 5.178mmol temp target (Client's list2 button); logs receipt of that TT alongside its own independent Tasker broadcast.

• AH<val> — AcceWeightHighDownTT/UpTT: remote ±0.01 nudge (via a 5.062/5.064mmol TT) on the boosted acce-ISF weight (ApsAutoIsfBgAccelWeightHigh); suffix is the new value.

• bat>1 — BatteryOver1%: phone battery recovered above 1% while on the battery-drop profile ("Current Profile50") → restores the standard profile and sets automation state Profile=AllOK.

• Batt1% — graph-announcement companion to Bt<1% (see below): fires in the same Battery1% block, rendered on the main graph via the announcement mechanism.

• BMild — BolusGivenMild: strengthens the SMB delivery ratio (tiered by BGL) on a mild delivery-driven BG rise without touching profile/iobTH, holding a brief 5.0mmol target for 2 minutes.

• BMildFS — BolusGivenMildFailsafe: safety-net sibling of BolusGivenMild for when SMBs simply haven't been delivering at all during a strongly-confirmed rise (near-zero IOB, 20-minute SMB silence) — same delivery-ratio boost response.

• Bol2 — "Bolus2": sets a brief 6.8mmol temp target after a bolus, either on a fast post-bolus BG fall (with COB/IOB present) or when carbs are up while MJ is active.

• BsUp — "BasalUp": raises acce weight back to neutral (0.50) and switches from the Low to the Standard profile once BGL is stable/rising outside an overnight/MJ guard window and activity is low.

• Bt<1% — Battery1%: phone battery ≤1% while on a normal running profile (PP130/C100/AllOK state) → switches to Current Profile50 and alerts (live pump only, not Virtual Pump).

• BTgOn / BTgOff — BoostToggleTT: manually setting a 5.08mmol TT remotely toggles ApsAutoIsfBoostAutomationsEnabled, the master switch for the BolusGiven/BolusGivenMild boost automations.

• CGrph — CleanGraphTT: manually setting a 5.042mmol TT remotely requests the "clean main graph" display combo (no SMB labels/arrowheads, solid green line) on the next OverviewFragment refresh.

• CLup — CloudLogsUploadTT: manually setting a 5.140mmol TT remotely triggers the same log upload as the Maintenance screen's "Send logs" button — for AAPSClient users with no direct settings access.

• COff1-N (N=1–2) — CarbsTHoff: restores iobTH to 70% (and acce to 0.50) when a post-carb BGL is falling or shows a mid-range anomaly; suffix identifies which of 2 condition blocks matched.

• Coff2-N (N=1–3) — CarbsStopTT5.7: cancels the 5.7mmol hypo-protection TT once carbs/IOB/BGL evidence shows recovery; suffix identifies which of 3 condition blocks matched.

• Coff4 — "carbsStopTT1ok4.4": cancels a very-low (≤4.4mmol) TT once carbs are active (COB≥10) and BGL is rising with low IOB, 10:00–22:00.

• D<val> — DuraWeightDownTT/UpTT: remote ±0.1 nudge (via a 5.022/5.024mmol TT) on the duraISF-weight baseline (ApsAutoIsfDuraWeightNormal); suffix is the new value, formatted to 3 trailing characters since this setting's range runs to 3.00.

• ______deadpod2hrs? — "OldPod" graph announcement: fires once a pod is >60h old AND BGL has been continuously high (raw Libre >11.0mmol OR AAPS BGL >10.0mmol) for 2+ hours — reads as pod failure, never touches dosing.

• DelOff — SMB-delivery-ratio reset: restores the resting SMB delivery ratio baseline once no boost TT is active (fires once as a boosted ratio drops back to baseline).

• DuraRsc — OvernightDuraRescue: a narrow, one-shot 02:00–04:00 switch back UP to the Standard profile when duraISF genuinely dominates a plateaued overnight high on the Low profile, with no stacking and no recent low.

• EvCap — EveningIobCeiling: caps iobTH at 45% from 22:00–00:00 whenever it's currently higher, independent of BGL direction (built after the 27 Jul 2026 evening SMB-stacking incident).

• Eve — "EveningTH": lowers acce weight to 0.45 and caps iobTH at 45% in the evening (and switches to the Low profile when a low is predicted, or unconditionally 22:00–06:00) while MJ is still clear.

• G2On / G2Off — Graph2ToggleTT: manually setting a 5.138mmol TT remotely toggles the "graph2" carb-model-curve display.

• G5On / G5Off — Graph5ToggleTT: manually setting a 5.142mmol TT remotely toggles graph5, a fixed clone of the main graph minus basal.

• Giv-N (N=1–3) — BolusGiven71_0.70: post-bolus response that boosts iobTH to 71%, restores acce to normal, and sets 110% profile for 2 minutes; suffix identifies which of 3 trigger blocks matched (manual-bolus rise, high-BGL-with-COB rise, or delivery-driven rise with no recent manual bolus/carbs).

• Gntl5 — GentleHypoRisk: escalates hypo protection from the 50%-prepare state (acce 0.07) to the deeper "Skittles" state (acce 0.02, iobTH 50%) using UKF-smoothed raw CGM signals; also fires a graph announcement ("________________Gentle5") for the same event.

• H4 — AlarmHypo2's second graph-announcement marker, fired alongside __________A4 in the same block (see below for AlarmHypo1's own "_____H4").

• _____H4 — AlarmHypo1's graph-announcement marker: logs a hypo-alarm event (slow decline, an unconditional emergency floor at 3.0mmol, or a steps-driven variant) that drops acce weight to 0.10 and marks a recent low.

• HardStackDelOff — forced SMB-delivery-ratio reduction (baseline−0.03, not a full reset) whenever genuine SMB stacking is detected (short average gap AND high count in the last 5 minutes), independent of TT state.

• HI<val> — HigherIsfRangeWeightDownTT/UpTT: remote ±0.1 nudge (via a 5.068/5.070mmol TT) on the higher-ISF-range weight (ApsAutoIsfHighBgWeight); suffix is the new value.

• HnAM — "HighNight00AM" — currently DISABLED in code (highNight00AmEnabled = false): would raise iobTH to 51%/acce to 0.50 and set a 4.2mmol TT to correct a genuine overnight high, but is turned off because it fights NightIobCeiling.

• IP<val> — PeakInsulinTimeDownTT/UpTT: remote ±5-minute nudge (via a 5.074/5.076mmol TT) on the configured insulin peak-activity time (InsulinOrefPeak); suffix is the new value in minutes, zero-padded to 3 digits.

• L<val> — LibreOffsetDownTT/UpTT: remote ±0.05 nudge (via a 5.032/5.034mmol TT) on the Libre calibration-offset baseline (ApsAutoIsfLibreOffsetOrig); suffix is the new value, allowed to run longer than the usual 5 characters since this setting is ≥1.0.

• L10Off — LowRaw24Cal ending: reverts the temporary Libre slope/offset override (0.75/1.3) once raw Libre BGL is no longer continuously below 10.0mmol for 24h.

• L10On — LowRaw24Cal starting: applies a temporary Libre slope/offset override (0.75/1.3) once raw Libre BGL has been continuously below 10.0mmol for 24h.

• LibreLd / Set1Ld — LibreVsSet1Race: logs which of the LibreSpecial/UKFset1 smoothing types is first to cross a -0.2mmol/5min drop threshold in a given episode (only the leader fires, once per episode); distinct 5-char prefixes keep them distinguishable under the graph's note-label truncation.

• LS<val> — LibreSlopeDownTT/UpTT: remote ±0.01 nudge (via a 5.026/5.028mmol TT) on the Libre calibration-slope baseline (ApsAutoIsfLibreSlopeOrig); suffix is the new value.

• MB<val> — MildBoostDownTT/UpTT: remote ±0.01 nudge (via a 5.052/5.054mmol TT) on the mild-boost SMB delivery ratio alone (ApsAutoIsfMildBoostRatio), distinct from the paired SB/SM nudge below; suffix is the new value.

• MJ — graph announcement fired alongside "MJ active" (below) when the direct MJ-cycle START user action runs; renders on the main graph via the TE.asAnnouncement mechanism instead of a plain note.

• MJ active — direct MJ-cycle START user action: begins an MJ (Mounjaro) dosing cycle — acce 0.35, iobTH 70%, switches to the Low profile — and sets automation state MJ=MJ active.

• MJ2 — old MJ-cycle port: advances MJ state from "MJ active" to MJ2 during 02:10–03:10.

• MJ3 — old MJ-cycle port: advances MJ state from MJ2 to MJ3 during 01:05–02:05.

• MJBOn / MJBOff — MjKotlinButtonsToggleTT: manually setting a 5.164mmol TT remotely toggles whether the Kotlin-ported MJ direct-action buttons are shown/enabled.

• MJoff-N (N=1–2) — old MJ-cycle port "MJoff": exits the MJ cycle from MJ3 back to NOMJremains, either on high BGL during the day (block 1, 12:00–21:04, BGL≥10.5mmol) or at midnight (block 2, 00:00–01:00).

• MJrec — "MJ recent or Steps CurrProf Acce": overnight (00:00–08:00) tuning applied while MJ is still active and the pod is fresh (<72h), setting acce 0.50/iobTH 70% and (HP2-gated) switching to the Low profile.

• MJs3 — MjStateMj3TT: manually setting a 5.146mmol TT forcibly sets the MJ automation state directly to MJ3 (manual override, paired with MJsNO as a "set it now" control).

• MJsNO — MjStateNoMjTT: manually setting a 5.144mmol TT forcibly sets the MJ automation state directly to NOMJremains (manual override).

• ML<val> — AutoIsfMaxLowDownTT/UpTT: remote ±0.1 nudge (via a 5.080/5.082mmol TT) on autoISF_max_low (ApsAutoIsfMaxLow — not yet wired into any calculation); suffix is the new value.

• MN<val> — AutoIsfMaxNormalDownTT/UpTT: remote ±0.1 nudge (via a 5.086/5.088mmol TT) on the live autoISF_max ceiling (ApsAutoIsfMax); suffix is the new value.

• MoreMJ — "MoreMJ": advances MJ state to MJ3 when acce weight is very low, activity is low, BGL is falling on the 50% profile with no carbs and IOB still present (or, on the OR-path, no genuine recent high has been seen while MJ still allows it).

• Night — "NightAcce_0.35TH18": overnight (01:00–06:00) reset of acce to 0.35 and iobTH to 18% (plus an automatic settings-export backup) when iobTH is out of the 18% band, no carbs are active, and BGL is low.

• No50 — "Not50%Recently": clears the LowBG=50recent flag once the profile is back at 100% and BGL is rising (delta ≥0.1mmol).

• NOMJremains — direct MJ-cycle RESTORE user action: ends the MJ dosing cycle (restores standard profile, acce 0.50, iobTH 70%), sets automation state MJ=NOMJremains.

• NtCap — NightIobCeiling: caps iobTH at 18% and acce at 0.35 from 00:00–06:00, independent of BGL, whenever either is currently looser (built after the 6→7 Aug 2026 hypo incident).

• off120 — "High6PPoff": exits the 120% profile boost once no TT exists at 120%, or BGL/delta have stabilised in the evening window at 120%.

• OffP-N (N=1–2) — "OffHighProf": overnight, when NOT on the Low profile and BGL is falling with a predicted low, drops to acce 0.18/iobTH 18% (and switches to the Low profile unless a rescue is active); suffix identifies which of 2 condition blocks matched.

• Old — "HighOldPod": sets a brief 5.0mmol TT + 110% profile boost when the cannula is stale (≥60h) or very fresh (≤6h) and BGL is high, rising, with a bolus ≥90 min ago.

• OldSensor<tier> — OldSensorAdj: tiered Libre slope/offset compensation for an aging (11–15 day) or very-new (0–3 day) sensor when raw Libre BGL has recently been >12.0mmol; tier suffix (NewDay1/NewDay2/NewDay3 for sensor days 0–3, or 1/2/3 for days 11–15) identifies which age band matched.

• OldSensorOff — OldSensorAdj reverting: restores the configured Libre slope/offset baseline once outside all the OldSensorAdj age tiers (or the latch was stale relative to the live values).

• P120 — "High6PP": brief 120% profile boost (3 condition blocks) when BGL is high with controlled delta and low/no TT tolerance, gated on a genuine recent high (raw Libre >12.0mmol/48h).

• PH<val> — PpWeightHighDownTT/UpTT: remote ±0.01 nudge (via a 5.056/5.058mmol TT) on the boosted ppISF weight (ApsAutoIsfPpWeightHigh); suffix is the new value.

• Pod130 — "PodChangeHighPP130": brief 130% profile boost around a pod change (fresh ≤2h or stale ≥78h) when BGL is high and rising, 10:00–18:00, gated on a genuine recent high.

• _____POD2 — "POD 78 hours": pod age reaches ~78 hours during 08:00–23:59 → switches to the Standard profile (Current ProfileReal), sets automation state Profile=PP130, and sends an SMS alert; unlike its sibling POD 79 h, this one is a real profile-switching action, not just a reminder (live pump only).

• POD 79 h — "POD 79 hours" reminder (graph announcement, no dosing action): pod age has reached ~79 hours, 07:00–23:59.

• PP<val> — PpWeightDownTT/UpTT: remote ±0.01 nudge (via a 5.012/5.014mmol TT) on the ppISF-weight baseline (ApsAutoIsfPpWeightNormal); suffix is the new value.

• PreSoak24hrs — "PreSoakSENSOR24hrs" reminder (graph announcement): reminds to pre-soak a new sensor at ~14.0 or ~14.5 days into the current sensor's life (cannula age permitting).

• PrfL — ProfileLowTT: manually setting a 5.150mmol TT remotely switches straight to the Low profile ("set it now", not a lock — other automations can still change it afterwards).

• PrfS — ProfileStandardTT: manually setting a 5.148mmol TT remotely switches straight to the Standard profile.

• PstRs — "PersistentRiseRelease": clears the way for normal dosing (cancels any active TT, forces the mild-offset-zero flag) once a genuinely sustained mild rise (delta ≥+0.10mmol, zero SMB delivered) has held continuously for ≥10 minutes.

• pTTOff — "RecentPodOff": relaxes acce weight back to baseline (and switches to the Standard profile) once a RecentPod/OldPod2 safety boost has ended and acce is still at its boosted level.

• RecPod — "RecentPod": brief 130% profile boost + 4.2mmol TT + boosted acce weight when the cannula is very fresh or stale, carbs are up, and BGL is rising with IOB headroom, gated on a genuine recent high (live pump only).

• SACOn / SACOff — SensorAgeCodeToggleTT: manually setting a 5.156mmol TT remotely toggles the master enable switch for the whole SensorAge Libre-calibration routine (distinct from STgOn/STgOff's narrower preference toggle).

• SB<val> — SmbDeliveryDownTT/UpTT: remote ±0.01 nudge (via a 5.002/5.004mmol TT) on the SMB delivery ratio baseline (ApsAutoIsfSmbDeliveryBaseline), written together with the paired SM nudge (see below) by the same TT; suffix is the new value.

• Semi — "SemiTwilightAcce_0.50TH16": drops acce weight to 0.50 and iobTH to 16% across 3 early-morning branches (sustained activity, very-high BGL with low acce, or a stable rising mid-morning window).

• SensorAgeCodeOff — OldSensorAdj master-disable path: restores the Libre baseline (and clears the active flag) if the whole SensorAge routine is disabled while an override was still active.

• Set50-N (N=1–4) — "prepare Set50%": drops into the 50%-profile hypo-protection state (acce 0.07, iobTH 50%, Low profile for 6h) on falling BGL; suffix identifies which of 4 condition blocks matched (dual-delta confirmation, exercise+IOB risk, native low-BGL fallback, or pre-sleep falling).

• _____S1hr — "SENSOR at 14.9 days" reminder (graph announcement): the current sensor is nearing its ~15-day replacement point (possible new-sensor overlap warning).

• _____S2hr — "SENSOR at 14 days 22 hours due" reminder (graph announcement): a slightly earlier sensor-age checkpoint than _____S1hr.

• Shwr12 — "Shower12": drops iobTH to 12% during a quiet morning rise (no steps, no COB, ≥180 min since last bolus) between 05:30–08:30.

• SM<val> — SmbDeliveryDownTT/UpTT (second value written alongside SB): remote ±0.01 nudge on ApsAutoIsfMildBoostRatio, made together with the SB baseline nudge by the same 5.002/5.004mmol TT.

• SO<val> — SmbOffsetDownTT/UpTT: remote ±0.1 nudge (via a 5.036/5.038mmol TT) on the SMB-offset override (ApsAutoIsfSmbOffsetOverride); suffix is the new value.

• _____ST601k — "Exercise limit Acce" alert (graph announcement): fires when a fast BG rise (≥0.4mmol/5min) coincides with very high recent step activity (≥1000 steps/60min).

• StBOn / StBOff — SteroidKotlinButtonToggleTT: manually setting a 5.166mmol TT remotely toggles whether the Kotlin-ported Steroid direct-action buttons are shown/enabled.

• Steps Steroids OFF — "Steps Steroids OFF": asserts Steroids=Steroids Off based on sustained activity (≥1000 steps/60min) with moderate IOB and controlled glucose — the only automated writer of the Steroids-off state; Steroids ON is always a manual action.

• Steroids130 — direct Steroid-button INCREASE_130 action: steps the steroid dosing ladder up to the 130% tier (iobTH 73%, acce 0.70, switches to Steroid Profile130).

• Steroids150 — direct Steroid-button INCREASE_150 action: steps the ladder up to 150% (iobTH 74%, switches to Steroid Profile150).

• Steroids190 — direct Steroid-button INCREASE_190 action: steps the ladder up to 190% (iobTH 75%, switches to Current Profile190Real).

• Steroids250 — direct Steroid-button INCREASE_250 action: steps the ladder up to 250% (iobTH 76%, switches to Current ProfileReal250).

• SteroidsOff — direct Steroid-button TURN_OFF action: exits the steroid dosing ladder, restores Current ProfileReal, iobTH 71%, acce 0.71, sets automation state Steroids=Steroids Off.

• SteroidsON — direct Steroid-button START_110 action: enables the steroid dosing ladder at its 110% tier (iobTH 71%, acce 0.70, switches to Steroid Profile110).

• STgOn / STgOff — SensorAgeToggleTT: manually setting a 5.006mmol TT remotely toggles the OldSensorAdj enable preference (ApsAutoIsfOldSensorAdjEnabled) — a narrower switch than SACOn/SACOff's master routine toggle.

• StLow <free text> — low-storage warning: the real-pump phone writes "StLow X.XGB free" as a plain CarePortal note once free space drops under 10GiB (re-arms above 10.5GiB); the Virtual Pump phone relays that same note as a graph announcement once it sees it via Nightscout (its relay logic also recognises a legacy "StorageLow " prefix, though nothing in this file currently writes that prefix itself).

• Sub75Cap — "Sub75HeavyDelivery": arms a 10-minute SMB cooldown once >1.0U has been delivered within a rolling 10 minutes while BGL is still under 7.5mmol.

• Sub75Clr — Sub75HeavyDelivery early release: clears that cooldown early if BGL has already recovered above 7.5mmol and is climbing fast (delta >0.7mmol).

• T<sign><val> — TodOffset*DownTT/UpTT (8 window pairs spanning T1–T8, the 8 fixed time-of-day windows 00-02h/02-04h/04-06h/06-09h/09-12h/12-18h/18-22h/22-00h): remote ±0.1 nudge on one of the 8 ApsAutoIsfTodOffset* preferences via its own dedicated TT (5.092–5.136mmol); the note text carries only the resulting signed value, not which of the 8 windows changed.

• _____Test3 — "Test3" validation port: exercises the state-check → notification → graph-announcement mechanics end-to-end via a test-only MJ=MJ6 automation state (no real dosing effect).

• THfloor — "iobTHDaytimeFloor": restores iobTH to 70% during the day (08:00–22:00) whenever it's stuck at ≤50% with no active low/falling BGL/TT reason to be there.

• TOD0 <reason> — automatic clearing of one or more time-of-day ISF offsets, once a sustained (≥10min) condition (MJ state change, HP2 crossing 4.0/6.0, or a Steps-state trigger) makes the existing offset direction no longer appropriate; the reason text names which trigger fired.

• TT3-<block> — "Skittles" hypo-risk: sets a 5.7mmol/180min safety temp target across 7 condition blocks (A–G) ranging from an emergency-floor fallback to graded IOB/delta-based fall detection.

• TT8.0off — "T8.0Off3ok": cancels the manually-set 8.0mmol "hyp" temp target once BGL is confidently rising (with an activity/steps gate) and low-activity confirmed.

• TToff-<which> — TT57Reversal (replaces TToff2-5/HypoTTOff1/TT5.8New2-3/mildConfirmed): cancels the 5.7mmol hypo-protection TT once one of several recovery patterns is confirmed; suffix (off1–off5, N2, N3, or Mld) identifies which branch matched — off1 is the early-catch "HypoTTOff1" pattern, off2–off5 are the graded TToff2-5 recovery levels, N2/N3 are the TT5.8New2/3 high-G/low-IOB catches, and Mld hands off cleanly to BolusGivenMild's own boost.

• TWi — "TwilightTH15Acce0.50": drops iobTH to 15% and acce to 0.50 during a quiet early morning (06:00–08:00, low steps, flat/falling BGL) while iobTH sits in a mid-range (17–39%) band.

• UamBst — "UAM Boost" Tier 3 notification: reports that DetermineBasalAutoISF's own SMB-boost logic fired this cycle. DetermineBasalAutoISF.kt is a pure calculation class with no I/O of its own — it only sets a flag (`uamBoostFiredThisCycle`); OpenAPSAutoISFPlugin.kt reads that flag back after calling `determine_basal()` and is the one place that actually writes the SMS/CarePortal note. Added 2026-08-22: also raises a graph-only "B" annotation via addGraphAnnouncement() alongside this note -- no additional SMS or alert, purely a graph marker.

• UKF1Block — LibreUkf1ToggleTT blocked path: on a real pump (not Virtual Pump, or on AAPSCLIENT), the UKFset1 live-comparison toggle is forced off and the manual toggle attempt is blocked rather than honoured.

• UKF1VOn / UKF1VOff — LibreUkf1ToggleTT: on Virtual Pump only (and not AAPSCLIENT), manually setting a 5.152mmol TT toggles the UKFset1 smoothing live-comparison feature (FslUseUkfSmoothing).

• UsuIP-N (N=1–3) — "Usual2forTH70": restores iobTH to 70% and acce to 0.50 once BGL has recovered (≥5.5mmol, stabilised, Steroids Off); suffix identifies which of 3 recovery blocks matched.

• W<val>% — WizardPctDownTT/UpTT: remote ±5 nudge (via a 5.046/5.048mmol TT) on the Wizard bolus percentage (OverviewBolusPercentage); suffix is the new percentage, zero-padded to 3 digits.

• X50-N (N=1–3) — "Extra50%": deepens hypo protection (acce 0.07, iobTH 50%) when BGL is falling on the 100% profile; suffix identifies which of 3 condition blocks matched (steep multi-delta fall, moderate fall, or overnight falling 01:00–05:00).

# Appendix E — List 1 and List 2 Temporary-Target Command Catalogue  [↑ TOC]

The mmol/L numbers below are command signals, not clinical glucose targets. On a pump or Virtual build, List 1 and the clinical List 2 actions invoke the matching local handler directly and create no temporary target. On AAPSClient, the same selection creates a five-minute relay TT because preferences/actions do not sync directly; the pump receives the code, performs the action, cancels the relay target and applies its relay repeat guard. Every command still passes through its confirmation popup. List 2 graph controls are always local display settings and never use relay TTs.

## List 1 - settings and command codes  [↑ TOC]

List 1 availability note. Code 5.152 is offered only on a non-Client Virtual-Pump comparison device. Code 5.154 is intentionally unused: UKF2 graph history is maintained without switching the live dosing engine, so checking a graph option no longer changes the glucose source.

## List 2 - actions, relay codes and local graph controls  [↑ TOC]

# Appendix F — Automation States In Use  [↑ TOC]

A reference for every named state this fork’s coded automations (and, for a few, native Automation-tab actions) read or write via the AutomationStateService — what it means, every automation that writes each value, and every automation that reads it as a gate. Generated by a full trace of every checkAutomationState()/setAutomationState()/automationStateService.* call site in OpenAPSAutoISFPlugin.kt, current as of the 2026-08-22 cleanup that removed three dead states (see §15g below).

## a. MJ  [↑ TOC]

Values: NOMJremains (default), MJ active, MJ2, MJ3, MJ4, MJ5, MJ6

Written by:

• handleDirectMjUserAction (Overview MJ button handler, not a readyToRun automation) — Action.START → "MJ active" (line 348); Action.RESTORE → "NOMJremains" (line 357)

• "MJ4" — MJ == MJ4 → "NOMJremains" (line 1922; port of the old "Test" automation)

• "MJ5" — MJ == MJ5 → "NOMJremains", also switches to Standard profile for 30 min (line 1930; port of "Test2")

• "Test3" — MJ == MJ6 → "NOMJremains" (line 1946; diagnostic port, also fires a URGENT notification + targeted SMS)

• "MJ2old" — MJ == MJ active AND time 02:10–03:10 → "MJ2" (line 1953)

• "MJ3old" — MJ == MJ2 AND time 01:05–02:05 → "MJ3" (line 1961)

• "MJoff" — MJ == MJ3 AND (12:00–21:04 with BG ≥10.5mmol, OR 00:00–01:00) → "NOMJremains" (line 1975)

• "MjStateNoMjTT" — manual remote-control TT code (≈5.144mmol) → "NOMJremains", unconditional (line 2799)

• "MjStateMj3TT" — manual remote-control TT code (≈5.146mmol) → "MJ3", unconditional (line 2807)

• "MoreMJ" — complex daytime condition tree (low acce ≤0.11, low activity, BG falling <5.5mmol on 50% profile, no COB, IOB≥0.2, LowBG=50recent, etc.) OR an independent "no recent genuine high" path → "MJ3" (line 4734)

Read by:

• handleDirectMjUserAction / handleDirectSteroidUserAction — re-validate the MJ/Steroid button press is still valid before acting (lines 326, 380, 390, 396, 402, 408)

• "MJ4"/"MJ5"/"Test3"/"MJ2old"/"MJ3old"/"MJoff" — each keys off the current MJ value to know when to advance/reset the cycle (lines 1920, 1927, 1940, 1951, 1959, 1968)

• TOD-offset auto-clear block — nomjRemains gates both the negative- and positive-offset clearing logic (line 2690)

• "Extra50" blocks 1/2 — require an MJ cycle to be *active* (!NOMJremains) as part of the fall-protection trigger (line 3423)

• "HighOldPod" — requires NOMJremains (line 3588)

• "HighNight00AM" branch 2 — requires NOMJremains (line 3659; block is permanently disabled via highNight00AmEnabled = false)

• "BolusGiven" bg2 branch — requires MJ NOT NOMJremains (an MJ cycle active) (line 3839)

• mjActive() helper (checks MJ == "MJ active" specifically) — blocks "BolusGiven" bg3 and "BolusGivenMild"'s mildConfirmed while MJ is freshly active (lines 3902, 4223)

• "MJrecentCurrProfAcce" — requires MJ NOT NOMJremains as one of several gates (line 4641)

• "BasalUp" — NOMJremains is one OR-branch of its cannulaOrStateOk gate (line 4670)

• "MoreMJ" — requires NOMJremains in both its main condition tree and its independent noRecentHighTrigger path (lines 4710, 4731)

• "AcceUp0.5" — requires NOMJremains OR SteroidsON (line 4854)

• "EveningIobCeiling" — requires NOMJremains (line 4974)

• "EveningTH" — evB1/evB2 require MJ NOT "MJ active"; evB3 requires "MJ active" (lines 5047, 5051, 5054)

• "Bolus2" branch b2 — requires "MJ active" (line 5213)

• autoIsfSettingsSnapshot() — logs the current MJ value into the AAPSClient settings snapshot (diagnostic only, line 5838)

Summary: MJ tracks the multi-stage state of a specific manual dosing protocol ("MJ 0.35/.70"), triggered by an Overview button press. NOMJremains is the neutral/default value; pressing the button advances it to "MJ active", and a chain of time-gated automations (MJ2old→MJ2, MJ3old→MJ3, MJoff→NOMJremains) walks it back down across the night, while the remaining numbered values (MJ4/MJ5/MJ6) are wired to test/diagnostic automations rather than the real cycle. The state is read constantly throughout the file as a gate: dozens of unrelated automations (evening/overnight profile, acceleration-weight and iobTH adjustments, post-bolus SMB boosts, TOD-offset clearing, basal-up recovery) branch on whether an MJ cycle is currently in progress, because active MJ dosing changes what "normal" dosing behaviour should look like.

## b. Steroids  [↑ TOC]

Values: Steroids Off (default), SteroidsON

Written by:

• handleDirectSteroidUserAction Action.START_110 (Overview steroid button) → "SteroidsON" (line 424)

• handleDirectSteroidUserAction Action.TURN_OFF (Overview steroid button) → "Steroids Off" (line 468)

• "StepsSteroidsOff" — sustained steps ≥1000 in the last hour, IOB 2.0–4.0U, BG ≤8.0mmol rising ≥0.5mmol, no COB → "Steroids Off" (line 5360). Per the source comment this is the *only* coded automation that ever asserts this state; turning steroids ON is always a manual user action with no automated counterpart.

Read by (all as a "steroids not currently active" precondition, unless noted):

• handleDirectSteroidUserAction's own self-checks for each escalation step 110→130→150→190→250→off (lines 381, 384, 389, 395, 401, 407)

• "SkittlesHypoRisk" block F (line 3350)

• "iobTHDaytimeFloor" (line 3402)

• "OffHighProf" (line 3513)

• "Shower12" (line 3614)

• "HighNight00AM" (disabled block) (line 3649)

• "Usual2forTH" (line 4469)

• "CarbsTHoff" (line 4517)

• "MJrecentCurrProfAcce" (line 4639)

• "AcceUp0.5" — OR'd with MJ NOMJremains (line 4854)

• "TwilightTH15Acce" (line 5104)

• "NightAcce" (line 5126)

• "SemiTwilightAcce" — all 3 branches (lines 5159, 5161, 5165)

Summary: Steroids records whether the user is currently on a steroid course, which materially raises insulin resistance and warrants the much higher iobTH/profile-percentage escalation ladder reachable via the manual 110%→130%→150%→190%→250% steroid buttons. Nearly every "back off"/hypo-protective automation elsewhere in the file (the Twilight/SemiTwilight/NightAcce resets, CarbsTHoff, Usual2forTH, OffHighProf, Shower12, etc.) gates on Steroids Off so it can't accidentally undo steroid-driven dosing; only one coded automation (StepsSteroidsOff) can clear it automatically, based on sustained-activity evidence that resistance has resolved.

## c. LowBG  [↑ TOC]

Values: 50recent, NO50rec (default)

Written by:

• "PrepareSet50" → 50recent (line 2015; deep-hypo entry: acce→0.07, iobTH→50, forces the Low profile)

• "PP50Off" → NO50rec (line 3306; 6-block recovery exit from the 50% prepare state)

• "SkittlesHypoRisk" → 50recent (line 3362; 7-block hypo-risk TT-5.7 set)

• "50SetRecent" → 50recent (line 3377; fires when profile actually dropped to 50% while state was still NO50rec)

• "Not50Recently" → NO50rec (line 3386; profile back at 100% and BG delta ≥0.1mmol -- as of 2026-08-23, also requires BG ≥5.5mmol; see the summary below)

• "Extra50" → 50recent (line 3435; deepens hypo protection on a confirmed fall)

• "TT57Reversal" off1–off5 branches → NO50rec (line 4233; the 5.7mmol TT genuinely ended/recovered)

• "CarbsStopTT57" → NO50rec (line 4450)

• "Usual2forTH" → NO50rec (line 4504)

• "CarbsTHoff" → NO50rec (line 4540)

• "HighPP130Off" → NO50rec (line 4829; also resets Profile to C100)

• "AlarmHypo1" → 50recent (line 5296; any of 4 hypo-alarm branches)

• "AlarmHypo2" → 50recent (line 5341; low-glucose + recent-steps alarm)

Read by:

• "PP50Off" gate — must be 50recent (line 3265)

• "50SetRecent" gate — must be NO50rec (plus profile==50) (line 3376)

• "Not50Recently" gate — must be 50recent (plus profile==100) (line 3384)

• "iobTHDaytimeFloor" gate — must be NO50rec (line 3401)

• "Extra50" gate — must be NO50rec (line 3418)

• "OvernightDuraRescue" — noRecentLow requires NOT 50recent, excluding rebound-highs from the rescue (line 3486)

• DetermineBasalAutoISF.kt's recent-low rebound guard — fed via the recentLowActive = checkAutomationState("LowBG","50recent") constructor param (line 5732)

• BolusWizard (external, per source comment) also reads this same flag to halve carb-bolus insulin after a recent low

• rt.reason export line, appended every cycle regardless of outcome: "LowBGrecent: Y/N" for the AIV history/CSV export (line 5807)

Summary: LowBG is the "was there a recent low" flag driving the file's 50%-profile hypo-protection state machine, and it also feeds two safety guards outside this state machine: the recent-low rebound guard in determine_basal() and BolusWizard's carb-bolus halving. It is set to 50recent whenever any of several hypo-risk/alarm automations fire (PrepareSet50, Skittles, Extra50, AlarmHypo1/2, 50SetRecent) and cleared back to NO50rec only once BG has genuinely recovered (PP50Off, Not50Recently, TT57Reversal, CarbsStopTT57, Usual2forTH, CarbsTHoff, HighPP130Off) — deliberately maintained as an actively-cleared two-way flag (unlike the removed BGLstate one-way latch), so downstream consumers can trust a "not recent" reading as much as a "recent" one.

The BG floor (≥5.5mmol) was chosen from real evidence: checked against all LowBG Y-to-N clear events across a real 4-day sample, clears below 5.5mmol were consistently followed by a large rebound correction, clears at or above 6.0mmol generally weren't. Since this flag gates the recent-low rebound guard (SMB halved while armed), keeping it armed longer means a correction landing during that window delivers at half strength rather than being skipped -- see the changelog for the full evidence writeup.

BG floor added 2026-08-23. Checked against all 20+ LowBG Y-to-N clear events across a real 4-day sample (20-23 Aug): the delta-only bar cleared the flag on the very first uptick off a low, often while BG was still clearly in low/low-normal range (4.2-5.0mmol seen repeatedly) -- not once the low had actually resolved. Clears below 5.5mmol were consistently followed by a large (≥0.8U) correction into a continuing rebound; clears at or above 6.0mmol generally weren't.

Worst confirmed case: 23 Aug, cleared at 4.9mmol/11:07am, an autonomous carbless rebound (COB=0.0 throughout) climbed unprotected to 7.6mmol by 12:41pm, ~1.2U was delivered, and that IOB decayed BG back down to 5.6mmol -- still falling -- three-plus hours later. Since this same flag gates DetermineBasalAutoISF.kt's recent-low rebound guard (halves any SMB while COB>0 or an unannounced-carb-impact pattern is detected), keeping it armed longer means a real correction landing during that window now delivers at roughly half strength instead of full, rather than skipping correction altogether.

## d. Steps  [↑ TOC]

Values: StepsLow, StepsHigh (no default declared)

Written by: none in this file. Read-only from this file — owned externally (native Automation-tab action or other integration), per the source's own bootstrap comment.

Read by:

• TOD-offset auto-clear block — stepsHigh alone can force clearNegativeTodOffsetsNow true (protective during activity-driven fall risk); stepsLow is one of several ANDed conditions (with NOMJremains and HP2 >6.0) required before clearPositiveTodOffsetsNow fires (lines 2714–2719). This is Steps' only use site in the file.

Confirmed misconfigured on the live phone (2026-08-22): the States tab there has no Steps state at all — instead there is a separate, self-named StepsLow state (state name StepsLow, holding value StepsLow), almost certainly from whatever native Automation-tab action writes step data being configured to set state StepsLow directly rather than state Steps with value StepsLow/StepsHigh. Since the code only ever checks state name Steps, checkAutomationState("Steps", "StepsLow") and ("Steps", "StepsHigh") both read false unconditionally on that device. Practical effect: clearPositiveTodOffsetsNow (which requires stepsLow with no alternative path) can never fire at all, so positive TOD offsets never auto-clear; clearNegativeTodOffsetsNow still works via its other two OR-branches (MJ state, HP2<4.0) but permanently loses the stepsHigh trigger path. Fix is phone-side: retarget the writing action to state=Steps, value=StepsLow, and add a StepsHigh-writing action if one doesn't already exist elsewhere.

Summary: Steps is a read-only external activity classification (StepsLow/StepsHigh) consumed solely by the time-of-day ISF-offset auto-clear logic: a high-activity reading alone is enough to force-clear negative TOD offsets, while a low-activity reading is one of several conditions required before positive TOD offsets are cleared.

## e. Profile  [↑ TOC]

Values: PP130, C100, AllOK (default), Batt1%, Bolus

Written by:

• "BatteryOver1pc" → AllOK (line 3579; battery recovered above 1%)

• "Pod2" (POD 78 hours) → PP130 (line 4598)

• "ExportSettingsPodActivation" → PP130 (line 4946; fires when a new pod has been activated since the last settings export)

• "PodChangeHighPP130" → C100 (line 4802; starts a 130% boost around a pod change)

• "HighPP130Off" → C100 (line 4826; exits the 130%/110% boost)

• "BolusGiven" (bg1/bg2/bg3 branches) → Bolus (line 3934; marks a post-bolus 6.8mmol TT so it isn't misread as the manual Activity TT)

• "Bolus2" → Bolus (line 5219; same marker, for its own post-bolus TT)

• Standalone unconditional block (no readyToRun key of its own) → C100 once >15min have passed since the last bolus, clearing the Bolus marker (line 5233)

• "HighNight00AM" (permanently disabled block, highNight00AmEnabled = false) → C100 (line 3665) — dead in practice

Read by:

• "Battery1pc" precondition — must currently be PP130, C100, or AllOK (a "normal running" state) before the battery-critical profile drop is allowed (line 3554)

• "ActivityProf50" — must NOT be Bolus, so a real Activity TT isn't confused with a brief post-bolus TT sharing the same 6.8mmol value (line 3681)

• "PodChangeHighPP130" — must be PP130 to start its own 130% boost (line 4795)

• The Bolus-clear block itself reads Bolus to know whether there's anything to clear (line 5230)

Note: Batt1% is declared as a valid value but is never actually written by any automation — Battery1pc's own source comment (line 3570) confirms this was a stale guard requirement that the writer never satisfied, since removed from the read side.

Summary: Profile is a catch-all marker state, mainly used to disambiguate *why* the effective profile-percentage is currently elevated (PP130/C100 from pod-change- or export-triggered boosts) and to flag a brief post-bolus temp target (Bolus) so it can't be confused with the manually-set Activity TT that happens to share its 6.8mmol value. AllOK is the neutral/default value.

## f. Sleeping  [↑ TOC]

Values: True (only value declared — there is no explicit "not sleeping" value; absence of a current value / a false inState read represents "not sleeping")

Written by: none in this file. Read-only from this file — owned externally (native Automation-tab trigger/action or other integration).

Read by:

• activityMonitor() — existSleepState = automationStateService.hasStateValues("Sleeping") and useSleepState = automationStateService.inState("Sleeping", "True") (lines 5984–5985), logged every cycle via aapsLogger.debug. useSleepState combined with recent steps ≤200 disables the inactivity-boost detection outright ("sleeping state", line 5999). existSleepState (whether the state has been declared at all, independent of its value) also gates the idle-window inactivity-detection branch via !existSleepState (line 6003).

Summary: Sleeping is a read-only external flag that the Activity Monitor's inactivity-sensitivity logic reads to suppress its own inactivity-boost detection while the user is asleep, since a stationary phone / low step count overnight should not be treated the same as daytime inactivity.

## g. Removed 2026-08-22  [↑ TOC]

Three states were removed from REQUIRED_AUTOMATION_STATES (and their remaining code references cleaned up) after an audit found no live use:

• MJstate (values MJon/MJoff) — every setState call for it was already commented out in the source; MJ (NOMJremains/MJ active/MJ2/MJ3/...) is the real, actively-used state. This looks like an earlier, abandoned naming attempt that was never fully cleaned up.

• BGLstate (value BGLlastLOW) — a write-only one-way latch, set in GentleHypoRisk/AlarmHypo1/AlarmHypo2 and cleared nowhere, so once any hypo alert had ever fired it stayed BGLlastLOW permanently and carried no information. Its one remaining reader (in MoreMJ) had already been swapped to the properly-maintained LowBG=50recent flag before this cleanup; the writes themselves were removed 2026-08-22 to make it genuinely inert rather than just functionally dead.

• query_got_up (value query_it) — a one-shot signal written in activityMonitor() whenever Sleeping=True and no steps were recorded between the 15-minute and 30-minute windows, but never read back by anything, here or elsewhere in the codebase. Sleeping itself remains live and is documented above.

# Appendix G — Current Live Settings Snapshot [↑ TOC]

Snapshot source: AutoISF_settings_Live_20260824_155639.txt
Captured: 24 August 2026, 3:56 PM — MAIN_PHONE_LOCAL (Live), full configuration, version 3.4.2.6+aisf321UK_654.

Important: These are point-in-time values captured from the Live device, not recommended defaults. Later app activity, profile changes, imports, or manual edits may change them.

| Setting | Live value |
| --- | --- |
| main_phone_snapshot_time | 8/24/26 03:56 PM |
| source | MAIN_PHONE_LOCAL |
| configuration_flavor | full |
| configuration_version | 3.4.2.6+aisf321UK_654 |
| configuration_application_id | info.nightscout.androidaps |
| configuration_profile | Current ProfileReal |
| Enable alternative activation of SMB always | true |
| autoISF_max | 2.60 |
| autoISF_max_low | 1.00 |
| autoISF_min | 0.30 |
| autoisf_bgaccel_isf_weight_high | 0.95 |
| autoisf_bgaccel_isf_weight_normal | 0.70 |
| autoisf_boost_automations_enabled | true |
| autoisf_clean_graph_requested | false |
| autoisf_custom_automations_enabled | true |
| autoisf_dura_isf_weight_normal | 2.50 |
| autoisf_fslcal_offset_normal | 1.40 |
| autoisf_fslcal_slope_normal | 0.72 |
| autoisf_libre_offset_orig | 1.40 |
| autoisf_libre_slope_orig | 0.72 |
| autoisf_low_profile_name | Current Profile |
| autoisf_low_raw_24_override_active | false |
| autoisf_low_storage_notified | false |
| autoisf_mild_boost_ratio | 0.21 |
| autoisf_mild_offset_zero_active | false |
| autoisf_mj_kotlin_buttons_enabled | true |
| autoisf_old_pod_notified | false |
| autoisf_old_sensor_adj_active | false |
| autoisf_old_sensor_adj_enabled | true |
| autoisf_pp_isf_weight_high | 0.15 |
| autoisf_pp_isf_weight_normal | 0.10 |
| autoisf_sensor_age_code_enabled | true |
| autoisf_smb_delivery_baseline | 0.13 |
| autoisf_smb_offset_override | 0.80 |
| autoisf_smb_offset_override_enabled | true |
| autoisf_standard_profile_name | Current ProfileReal |
| autoisf_steroid_kotlin_button_enabled | true |
| autoisf_tdd_factor | true |
| autoisf_tdd_factor_fallback | 1.00 |
| autoisf_tdd_sensitivity | true |
| autoisf_tod_offset_0002 | 0.00 |
| autoisf_tod_offset_0204 | 0.00 |
| autoisf_tod_offset_0406 | 0.00 |
| autoisf_tod_offset_0609 | 0.00 |
| autoisf_tod_offset_0912 | 0.00 |
| autoisf_tod_offset_1218 | 0.00 |
| autoisf_tod_offset_1822 | 0.00 |
| autoisf_tod_offset_2200 | 0.00 |
| autoisf_uam_boost_enabled | false |
| bgAccel_ISF_weight | 0.70 |
| bgBrake_ISF_weight | 0.25 |
| boost_bolus_cap | 2.50 |
| boost_max_iob | 1.00 |
| boost_scale_value | 1.00 |
| dura_ISF_weight | 2.50 |
| fslCal_Offset | 1.40 |
| fslCal_Slope | 0.72 |
| graph5_bgl_only | false |
| half_basal_exercise_target | 6.90 |
| high_temptarget_raises_sensitivity | true |
| higher_ISFrange_weight | 0.80 |
| iob_threshold_percent | 70 |
| low_temptarget_lowers_sensitivity | true |
| lower_ISFrange_weight | 0.70 |
| openapsama_enable_autoISF | true |
| openapsama_smb_delivery_ratio | 0.13 |
| openapsama_smb_delivery_ratio_bg_range | 0.00 |
| openapsama_smb_delivery_ratio_max | 0.50 |
| openapsama_smb_delivery_ratio_min | 0.95 |
| openapsama_smb_max_range_extension | 4.00 |
| pp_ISF_weight | 0.10 |
| show_carb_model_curve | true |
| show_graph5 | true |
| split_bolus_enabled | true |
| split_bolus_interval | 7 |

# End   [↑ TOC]